## Sheet1
Hazard ID | Hazard Description | Potential Harm | Severity (1-5) | Probability (1-5) | Initial Risk Level | Risk Control Measures | Residual Risk Level | Verification Method
HZ-001 | Incorrect medication dosage due to data entry error | Overdose or underdose leading to adverse effects | 5 | 3 | High | Input validation, dose calculation automation, confirmation dialogs | Medium | Unit and integration tests, usability tests
HZ-002 | Failure to alert on drug-drug interaction | Unintended drug reactions or toxicity | 4 | 2 | Medium | Automated interaction checker, clinical rule engine | Low | Rule engine test coverage, scenario simulations
HZ-003 | Incorrect TPN calculation based on outdated patient weight | Nutritional imbalance affecting patient recovery | 4 | 3 | High | Regular weight updates, TPN auto-adjust algorithms | Medium | Simulation with real-world data, clinical validation
HZ-004 | System unavailability during prescription entry | Delay in treatment, increased clinical risk | 3 | 4 | High | Redundant system architecture, offline mode for emergency use | Medium | Load testing, system failover tests